## QSP2350 cal 70783 8-6-20
|  | Calibration Date: | 2020-08-06 |  |  |  |  | Job No.: | L20183
|  | Model Number: | QSP2350
|  | Serial Number: | 70783
|  | Operator: | TPC |  | .
|  | Standard Lamp: | V-040(1/3/2019)
|  | Operating Voltage Range: | 6 | to | 15 | VDC (+)
Note: The QSP2350 output is a voltage that is proportional to the log of the incident irradiance.
To calculate irradiance, use this formula:
Irradiance = Calibration factor * (10^Light Signal Voltage - 10^Dark Voltage)
|  | Dry Calibration Factor: | 3.7641e+12 | quanta/cm²·sec per volt |  |  | 6.2506e-06 | µEinsteins/cm²·sec per volt
|  | Wet Calibration Factor: | 6.6457e+12 | quanta/cm²·sec per volt |  |  | 1.1036e-05 | µEinsteins/cm²·sec per volt
Sensor Test Data and Results2)
|  |  | Sensor Supply Current (Dark): | 3.47 | mA
|  |  | Supply Voltage: | 6 | Volts
|  |  | Lamp Integrated PAR Irradiance: | 9401310000000000 | quanta/cm2·sec |  | 0.015612 | µEinsteins/cm²sec
|  |  | Immersion Coefficient: | 0.5664
| Nominal Filter OD | Expected Transmission | Calibrated Trans. | Sensor Voltage | Expected Voltage | Voltage % Error | Measured Trans. | Transmission Error (%) | Test Irrad. (quanta/ cm2·sec)
5198.9676
9835.4501
| No Filter | 100% | 100% | 3.3977 | 3.3977 | 0% | 100% | 0 | 9401329198974062
| 0.3 | 50.12% | 36.1% | 2.946 | 2.9552 | -0.3125% | 35.32% | 2.2179 | 3.3203e+15
| 0.5 | 31.62% | 27.6% | 2.8338 | 2.8386 | -0.1697% | 27.27% | 1.222 | 2.5635e+15
| 1 | 10% | 9.27% | 2.358 | 2.3648 | -0.2875% | 9.09% | 1.982 | 854583308627346
| 2 | 1% | 1.11% | 1.422 | 1.443 | -1.478% | 1.018% | 9.0662 | 9.5699e+13
| 3 | 0.1% | 0.0534% | 0.228 | 0.12524 | 45.07% | 0.02744% | 94.6078 | 2.5989e+12
| RG780 | 0% | 0% | 0.19 | 0.0022095 | 98.84% | 0.02177% | -100 | 2.0658e+12
|  | Dark Before: | 0.00221 | Volts
|  | Light - No Filter Hldr.: | 3.3977 | Volts
|  | Dark After - NFH: | 0.002209 | Volts
|  | Average Dark | 0.0022095 | Volts
Notes:
1. Annual calibration is recommended.
2) This section is for internal use and for more advanced analysis.