## QSP2350 cal 70784 8-6-20
|  | Calibration Date: | 2020-08-06 |  |  |  |  | Job No.: | L20183
|  | Model Number: | QSP2350
|  | Serial Number: | 70784
|  | Operator: | TPC |  | .
|  | Standard Lamp: | V-040(1/3/2019)
|  | Operating Voltage Range: | 6 | to | 15 | VDC (+)
Note: The QSP2350 output is a voltage that is proportional to the log of the incident irradiance.
To calculate irradiance, use this formula:
Irradiance = Calibration factor * (10^Light Signal Voltage - 10^Dark Voltage)
|  | Dry Calibration Factor: | 3.7789e+12 | quanta/cm²·sec per volt |  |  | 6.2752e-06 | µEinsteins/cm²·sec per volt
|  | Wet Calibration Factor: | 6.6718e+12 | quanta/cm²·sec per volt |  |  | 1.1079e-05 | µEinsteins/cm²·sec per volt
Sensor Test Data and Results2)
|  |  | Sensor Supply Current (Dark): | 3.44 | mA
|  |  | Supply Voltage: | 6 | Volts
|  |  | Lamp Integrated PAR Irradiance: | 9401310000000000 | quanta/cm2·sec |  | 0.015612 | µEinsteins/cm²sec
|  |  | Immersion Coefficient: | 0.5664
| Nominal Filter OD | Expected Transmission | Calibrated Trans. | Sensor Voltage | Expected Voltage | Voltage % Error | Measured Trans. | Transmission Error (%) | Test Irrad. (quanta/ cm2·sec)
5219.3985
9874.1033
| No Filter | 100% | 100% | 3.396 | 3.396 | 0% | 100% | 0 | 9401388626539158
| 0.3 | 50.12% | 36.1% | 2.948 | 2.9535 | -0.1868% | 35.62% | 1.3512 | 3348703718533485
| 0.5 | 31.62% | 27.6% | 2.8356 | 2.8369 | -0.04617% | 27.49% | 0.41041 | 2.5842e+15
| 1 | 10% | 9.27% | 2.3576 | 2.3631 | -0.2324% | 9.116% | 1.6838 | 8.5715e+14
| 2 | 1% | 1.11% | 1.4269 | 1.4413 | -1.011% | 1.033% | 7.4383 | 9.7208e+13
| 3 | 0.1% | 0.0534% | 0.21 | 0.12354 | 41.17% | 0.02416% | 121.0478 | 2.3498e+12
| RG780 | 0% | 0% | 0.2 | 0.0089435 | 95.53% | 0.02267% | -100 | 2.2103e+12
|  | Dark Before: | 0.008942 | Volts
|  | Light - No Filter Hldr.: | 3.395 | Volts
|  | Dark After - NFH: | 0.008945 | Volts
|  | Average Dark | 0.0089435 | Volts
Notes:
1. Annual calibration is recommended.
2) This section is for internal use and for more advanced analysis.