## QSP2350 cal 70785 8-6-20
|  | Calibration Date: | 2020-08-06 |  |  |  |  | Job No.: | L20184
|  | Model Number: | QSP2350
|  | Serial Number: | 70785
|  | Operator: | TPC |  | .
|  | Standard Lamp: | V-040(1/3/2019)
|  | Operating Voltage Range: | 6 | to | 15 | VDC (+)
Note: The QSP2350 output is a voltage that is proportional to the log of the incident irradiance.
To calculate irradiance, use this formula:
Irradiance = Calibration factor * (10^Light Signal Voltage - 10^Dark Voltage)
|  | Dry Calibration Factor: | 3.8139e+12 | quanta/cm²·sec per volt |  |  | 6.3332e-06 | µEinsteins/cm²·sec per volt
|  | Wet Calibration Factor: | 6.7335e+12 | quanta/cm²·sec per volt |  |  | 1.1182e-05 | µEinsteins/cm²·sec per volt
Sensor Test Data and Results2)
|  |  | Sensor Supply Current (Dark): | 3.42 | mA
|  |  | Supply Voltage: | 6 | Volts
|  |  | Lamp Integrated PAR Irradiance: | 9401310000000000 | quanta/cm2·sec |  | 0.015612 | µEinsteins/cm²sec
|  |  | Immersion Coefficient: | 0.5664
| Nominal Filter OD | Expected Transmission | Calibrated Trans. | Sensor Voltage | Expected Voltage | Voltage % Error | Measured Trans. | Transmission Error (%) | Test Irrad. (quanta/ cm2·sec)
5267.6837
9965.4481
| No Filter | 100% | 100% | 3.392 | 3.392 | 0% | 100% | 0 | 9401335574413102
| 0.3 | 50.12% | 36.1% | 2.9425 | 2.9495 | -0.2381% | 35.5% | 1.7019 | 3337103241334211
| 0.5 | 31.62% | 27.6% | 2.831 | 2.8329 | -0.06743% | 27.45% | 0.54892 | 2.5806e+15
| 1 | 10% | 9.27% | 2.357 | 2.3591 | -0.08824% | 9.189% | 0.88541 | 863878346027579
| 2 | 1% | 1.11% | 1.4368 | 1.4373 | -0.0364% | 1.068% | 3.9056 | 1.0046e+14
| 3 | 0.1% | 0.0534% | 0.251 | 0.11954 | 52.37% | 0.03147% | 69.7013 | 2.9839e+12
| RG780 | 0% | 0% | 0.0029 | 0.0029025 | -0.08621% | -2.351e-07% | -100 | 25552311555.2767
|  | Dark Before: | 0.00291 | Volts
|  | Light - No Filter Hldr.: | 3.392 | Volts
|  | Dark After - NFH: | 0.002895 | Volts
|  | Average Dark | 0.0029025 | Volts
Notes:
1. Annual calibration is recommended.
2) This section is for internal use and for more advanced analysis.